## PM Express Menus
Menu 7-Seg | Page 7-Seg | Menu | Page | Menu Visible | Menu Security | Instance | Instance Visible | Class ID | Instance ID | Member ID | Class ID Name | Default Value 7-Seg | Parameter Value 7-Seg | Member ID Name | Parameter ID
|  | Analog Input | Operations | always | 4 | 1 | always | 0x04 | 255 | 0x01 | AIN |  |  | Value | 4001
|  | Analog Input | Operations |  |  | 1 |  | 0x04 | 255 | 0x02 | AIN |  |  | Error Status | 4002
|  | Analog Input | Operations |  |  | 1 |  | 0x04 | 255 | 0x0C | AIN |  |  | Calibration Offset | 4012
|  | Limit | Operations | If 3rd digit of model number is an L | 2 | 1 | always | 0x0C | 255 | 0x03 | LIM |  |  | Low Limit Set Point | 12003
|  | Limit | Operations |  |  | 1 |  | 0x0C | 255 | 0x04 | LIM |  |  | High Limit Set Point | 12004
|  | Control Loop | Operations | If 3rd digit of model number is a C | 4 | 1 | always | 0x08 | 255 | 0x01 | CLP |  |  | Control Mode | 8001
|  | Control Loop | Operations |  |  | 1 |  | 0x08 | 255 | 0x1A | CLP |  |  | Autotune | 8026
|  | Control Loop | Operations |  |  | 1 |  | 0x07 | 255 | 0x01 | SET |  |  | Set Point | 7001
|  | Control Loop | Operations |  |  | 1 |  | 0x08 | 255 | 0x09 | CLP |  |  | Heat Proportional Band | 8009
|  | Control Loop | Operations |  |  | 1 |  | 0x08 | 255 | 0x0A | CLP |  |  | On/Off Heat Hysteresis | 8010
|  | Control Loop | Operations |  |  | 1 |  | 0x08 | 255 | 0x0C | CLP |  |  | Cool Proportional Band | 8012
|  | Control Loop | Operations |  |  | 1 |  | 0x08 | 255 | 0x06 | CLP |  |  | Time Integral | 8006
|  | Control Loop | Operations |  |  | 1 |  | 0x08 | 255 | 0x07 | CLP |  |  | Time Derivative | 8007
|  | Control Loop | Operations |  |  | 1 |  | 0x07 | 255 | 0x02 | SET |  |  | Manual Power | 7002
|  | Alarm | Operations | always | 3 | 1 | always | 0x09 | 255 | 0x02 | ALM |  |  | Low Set Point | 9002
|  | Alarm | Operations |  |  | 1 |  | 0x09 | 255 | 0x01 | ALM |  |  | High Set Point | 9001
|  | Analog Input | Express Setup | always | 5 | 1 | always | 0x04 | 255 | 0x05 | AIN |  |  | Sensor Type | 4005
|  | Analog Input | Setup |  |  | 1 |  | 0x04 | 255 | 0x06 | AIN |  |  | TC Linearization | 4006
|  | Analog Input | Setup |  |  | 1 |  | 0x04 | 255 | 0x11 | AIN |  |  | Scale Low | 4017
|  | Analog Input | Setup |  |  | 1 |  | 0x04 | 255 | 0x12 | AIN |  |  | Scale High | 4018
|  | Analog Input | Setup |  |  | 1 |  | 0x04 | 255 | 0x14 | AIN |  |  | Display Precision | 4020
|  | Limit | Setup | If 3rd digit of model number is an L | 5 | 1 | always | 0x0C | 255 | 0x05 | LIM |  |  | Sides | 12005
|  | Limit | Setup |  |  | 1 |  | 0x0C | 255 | 0x02 | LIM |  |  | Hysteresis | 12002
|  | Control Loop | Setup | If 3rd digit of model number is a C | 5 | 1 | always | 0x08 | 255 | 0x03 | CLP |  |  | Heat Algorithm | 8003
|  | Control Loop | Setup |  |  | 1 |  | 0x08 | 255 | 0x04 | CLP |  |  | Cool Algorithm | 8004
|  | Control Loop | Setup |  |  | 1 |  | 0x07 | 255 | 0x0E | SET |  |  | Ramp Action | 7014
|  | Control Loop | Setup |  |  | 1 |  | 0x07 | 255 | 0x11 | SET |  |  | Ramp Rate | 7017
|  | Output | Setup | If digits 5 or 6 | 5 | 1 - 2 | Instance 1 if 5th digit is not an A Instance 2 if 6th digit is not an A | 0x06 | 255 | 0x05 | DIO |  |  | Output Function (output digital) | 6005
|  | Output | Setup |  |  | 1 - 2 |  | 0x06 | 255 | 0x03 | DIO |  |  | Fixed Time Base | 6003
|  | Output | Setup |  |  | 1 - 2 |  | 0x06 | 255 | 0x0A | DIO |  |  | High Power Scale | 6010
|  | Output | Setup |  |  | 1 |  | 0x12 | 255 | 0x01 | AOT |  |  | Output Type (output process) | 18001
|  | Output | Setup |  |  | 1 |  | 0x12 | 255 | 0x02 | AOT |  |  | Output Function | 18002
|  | Output | Setup |  |  | 1 |  | 0x12 | 255 | 0x09 | AOT |  |  | Scale Low | 18009
|  | Output | Setup |  |  | 1 |  | 0x12 | 255 | 0x0A | AOT |  |  | Scale High | 18010
|  | Alarm | Setup | always | 5 | 1 | always | 0x09 | 255 | 0x0F | ALM |  |  | Type | 9015
|  | Alarm | Setup |  |  | 1 |  | 0x09 | 255 | 0x03 | ALM |  |  | Hysteresis | 9003
|  | Alarm | Setup |  |  | 1 |  | 0x09 | 255 | 0x05 | ALM |  |  | Logic | 9005
|  | Alarm | Setup |  |  | 1 |  | 0x09 | 255 | 0x07 | ALM |  |  | Latching | 9007
|  | Alarm | Setup |  |  | 1 |  | 0x09 | 255 | 0x08 | ALM |  |  | Blocking | 9008
|  | Alarm | Setup |  |  | 1 |  | 0x09 | 255 | 0x06 | ALM |  |  | Silencing | 9006
|  | Alarm | Setup |  |  | 1 |  | 0x09 | 255 | 0x10 | ALM |  |  | Alarm Display | 9016
|  | Global | Setup | always | 5 | 1 | always | 0x03 | 1 | 0x05 | DSP |  |  | Display Units | 3005
|  | Communications | Setup |  | 5 | 1 |  | 0x11 | 255 | 0x01 | COM |  |  | Standard Bus Address | 17001
|  | Custom | Factory | always | 5 | 1 - 2 | always | 0x0E | 255 | 0x05 | CSU |  |  | Parameter | 14005
|  | Security | Factory | always | 1 | 1 | always | 0x03 | 1 | 0x0B | DSP |  |  | Set Lockout Security | 3011
|  | Diagnostics | Factory | always | 2 | 1 | always | 0x01 | 1 | 0x09 | DEV |  |  | Part Number | 1009
|  | Diagnostics | Factory |  |  | 1 |  | 0x01 | 1 | 0x11 | DEV |  |  | Software Revision | 1017
|  | Diagnostics | Factory |  |  | 1 |  | 0x01 | 1 | 0x05 | DEV |  |  | Software Build | 1005
|  | Diagnostics | Factory |  |  | 1 |  | 0x01 | 1 | 0x20 | DEV |  |  | Serial Number | 1032
|  | Diagnostics | Factory |  |  | 1 |  | 0x01 | 1 | 0x08 | DEV |  |  | Date of Manufacture | 1008
|  | Diagnostics | Factory |  |  | 1 |  | 0x01 | 1 | 0x0D | DEV |  |  | Restore Settings From | 1013
|  | Diagnostics | Factory |  |  | 1 |  | 0x03 | 1 | 0x1A | DSP |  |  | Zone | 3026
|  | Calibration | Factory | always | 5 | 1 | always | 0x04 | 255 | 0x15 | AIN |  |  | Electrical Measurement | 4021
|  | Calibration | Factory |  |  | 1 |  | 0x04 | 255 | 0x0A | AIN |  |  | Electrical Input Offset | 4010
|  | Calibration | Factory |  |  | 1 |  | 0x04 | 255 | 0x0B | AIN |  |  | Electrical Input Slope | 4011
|  | Calibration | Factory |  |  | 1 | If 5th digit is F | 0x12 | 255 | 0x05 | AOT |  |  | Part Number | 18005
|  | Calibration | Factory |  |  | 1 |  | 0x12 | 255 | 0x06 | AOT |  |  | Public Key | 18006
|  | Calibration | Factory | always | 5 | 1 | always | 0x01 | 1 | 0x2E | DEV |  |  | Electrical Output Offset | 1046
|  | Calibration | Factory | always | 5 | 1 | always | 0x01 | 1 | 0x2F | DEV |  |  | Electrical Output Slope | 1047
|  |  | Express Setup | always | 1 |  |  | 0x03 | 1 | 0x0B | DSP | 5 | LoC | Lockout Menu | 3011
|  |  | Express Setup |  |  |  |  | 0x04 | 1 | 0x05 | AIN | tC | SEn | Sensor Type | 4005
|  |  | Express Setup |  |  |  |  | 0x04 | 1 | 0x06 | AIN | J | Lnr | Linearization | 4006
|  |  | Express Setup |  |  |  |  | 0x04 | 1 | 0x14 | AIN | 0 | dEC | Decimal | 4020
|  |  | Express Setup |  |  |  |  | 0x03 | 1 | 0x05 | DSP | F | C_F | Display Units | 3005
|  |  | Express Setup |  |  |  |  | 0x04 | 1 | 0x11 | AIN | 0 | rLo | Range Low | 4017
|  |  | Express Setup |  |  |  |  | 0x04 | 1 | 0x12 | AIN | 9999 | rhi | Range High | 4018
|  |  | Express Setup |  |  |  |  | 0x06 | 1 | 0x05 | DIO | hEAt | Fn1 | Function of Output 1 (digital) | 6005
|  |  | Express Setup |  |  |  |  | 0x12 | 1 | 0x02 | AOT | hEAt | Fn2 | Function of Output 1 (analog) | 18002
|  |  | Express Setup |  |  |  |  | 0x12 | 1 | 0x01 | AOT | uoLt | o.ty | Output Type | 18001
|  |  | Express Setup |  |  |  |  | 0x06 | 2 | 0x05 | DIO | oFF | Fn2 | Function of Output 2 (digital) | 6005
|  |  | Express Setup |  |  |  |  | 0x08 | 1 | 0x03 | CLP | Pid | h.Ag | Heat Algorithm | 8003
|  |  | Express Setup |  |  |  |  | 0x08 | 1 | 0x04 | CLP | oFF | C.Ag | Cool Algorithm | 8004
|  |  | Express Setup |  |  |  |  | 0x08 | 1 | 0x0A | CLP | 3.0 | HSC | Hysteresis (Heat & Cool) | 8010
|  |  | Express Setup |  |  |  |  | 0x0C | 1 | 0x05 | LIM | both | L.Sd | Limit Sides | 12005
|  |  | Express Setup |  |  |  |  | 0x0C | 1 | 0x02 | LIM | 3.0 | L.hy | Limit Hysteresis | 12002
|  |  | Express Setup |  |  |  |  | 0x09 | 1 | 0x0F | ALM | oFF | A.ty | Alarm Type | 9015
|  |  | Express Setup |  |  |  |  | 0x09 | 1 | 0x03 | ALM | 1.0 | A.hy | Alarm Hysteresis | 9003
|  |  | Express Setup |  |  |  |  | 0x09 | 1 | 0x05 | ALM | A.LC | A.Lg | Alarm Logic | 9005
|  |  | Express Setup |  |  |  |  | 0x09 | 1 | 0x07 | ALM | nLAt | A.LA | Alarm Latching | 9007
|  |  | Express Setup |  |  |  |  | 0x09 | 1 | 0x08 | ALM | oFF | A.bL | Alarm Blocking | 9008
|  |  | Express Setup |  |  |  |  | 0x09 | 1 | 0x06 | ALM | oFF | A.Si | Alarm Silencing | 9006
|  |  | Express Setup |  |  |  |  | 0x09 | 1 | 0x10 | ALM | oFF | A.dSP | Alarm Display | 9016
|  |  | Express Setup |  |  |  |  | 0x07 | 1 | 0x0E | SET | oFF | rP | Ramp Action | 7014
|  |  | Express Setup |  |  |  |  | 0x07 | 1 | 0x11 | SET | 1.0 | r.rt | Ramp Rate | 7017
|  |  | Express Setup |  |  |  |  | 0x06 | 1 | 0x0A | DIO | 100.0 | o.hi1 | Power Scale High Output 1 | 6010
|  |  | Express Setup |  |  |  |  | 0x12 | 1 | 0x09 | AOT | 0.0 | S.Lo | Scale Low Output 1 | 18009
|  |  | Express Setup |  |  |  |  | 0x12 | 1 | 0x0A | AOT | 10.0 | S.hi | Scale High Output 1 | 18010
|  |  | Express Setup |  |  |  |  | 0x06 | 2 | 0x0A | DIO | 100.0 | o.hi2 | Power Scale High Output 2 | 6010
|  |  | Express Setup |  |  |  |  | 0x0E | 1 | 0x05 | CSU | AC.Pu | PAr1 | Upper or Left Display | 14005
|  |  | Express Setup |  |  |  |  | 0x0E | 2 | 0x05 | CSU | AC.SP | PAr2 | Lower or Right Display | 14005
|  |  | Express Setup |  |  |  |  | 0x11 | 255 | 0x01 | COM | 1 | Ad.S | Zone Address | 17001